ANALYTICAL METHOD DEVELOPMENT AND VALIDATION FOR SIMULTANEOUS DETERMINATION OF SUMATRIPTAN AND NAPROXEN BY RP-HPLC

TABLES AND FIGURES:

Table -1

Linearity:-

Linearity data of Naproxen

| S.No | Linearity Level | Concentration | Area |
| --- | --- | --- | --- |
| 1 | I | 60ppm | 1189032 |
| 2 | II | 70ppm | 1497165 |
| 3 | III | 80ppm | 1851004 |
| 4 | IV | 90ppm | 2186380 |
| 5 | V | 100ppm | 2548658 |
| Correlation Coefficient |  |  | 0.999 |

Table -2

Linearity data of Sumatriptan:-

| S.No | Linearity Level | Concentration | Area |
| --- | --- | --- | --- |
| 1 | I | 60ppm | 1219297 |
| 2 | II | 70ppm | 1531332 |
| 3 | III | 80ppm | 1856868 |
| 4 | IV | 90ppm | 2252060 |
| 5 | V | 100ppm | 2560402 |
| Correlation Coefficient |  |  | 0.999 |

| Parmeters | Results observed Sumatriptan | Results observed Naproxen |
| --- | --- | --- |
| Slope | 34029.38 | 34898.55 |
| Intercept | -838359 | -945575 |
| Correlation | 0.999168 | 0.999901 |

Table -3

Linearity parameters

Table -4

Assay:-Sumatriptan and Naproxen

| S.no | Naproxen |  |  |  | Sumatriptan |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Standard Rt | Sample Rt | Standard area | Sample area | Standard Rt | Sample Rt | Standard area | Sample area |
| 01 | 2.248 | 2.24 | 1523510 | 1521502 | 5.871 | 5.865 | 1566643 | 1561262 |
| 02 | 2.249 | 2.250 | 1531670 | 1529265 | 5.875 | 5.863 | 1573899 | 1564536 |
| MEAN |  |  | 1532594 | 1525384 |  |  | 1570271 | 1562899 |
| STD |  |  | 1527590 | 5488.8 |  |  | 1570271 | 2315.2 |
| %RSD |  |  | 0.377 | 0.36 |  |  | 0.327 | 0.15 |

Table -5

Accuracy:-

Accuracy data of Sumatriptan

| %Concentration (at specification Level) | Area | Amount Added (mg) | Amount Found (mg) | % Recovery | Mean Recovery |
| --- | --- | --- | --- | --- | --- |
| 50% | 984243 | 6.15 | 6.23 | 101.4% | 101.0% |
| 100% | 1567396 | 10.0 | 9.98 | 99.8 |  |
| 150% | 2497228 | 15.5 | 15.79 | 101.9% |  |

Table -6

Accuracy data of Naproxen:-

| %Concentration (at specification Level) | Area | Amount Added (mg) | Amount Found (mg) | % Recovery | Mean Recovery |
| --- | --- | --- | --- | --- | --- |
| 50% | 958196 | 6.15 | 6.25 | 101.7% | 101.0% |
| 100% | 1532695 | 10.0 | 9.96 | 99.5% |  |
| 150% | 2425792 | 15.6 | 15.8 | 101.7% |  |

Table -7

Precision:

| SI.No | Injection number (80 mcg/ml) | Retention Time of Naproxen | Retention Time of Sumatriptan | Area of Naproxen | Area of Sumatriptan |
| --- | --- | --- | --- | --- | --- |
| 1 | Injection-1 | 2.412 | 6.374 | 1549491 | 1579612 |
| 2 | Injection-2 | 2.259 | 5.900 | 1530248 | 1574521 |
| 3 | Injection-3 | 2.259 | 5.884 | 1530713 | 1566564 |
| 4 | Injection-4 | 2.261 | 5.863 | 1527834 | 1570849 |
| 5 | Injection-5 | 2.257 | 5.857 | 1537667 | 1567180 |
|  | AVG |  |  | 1535191 | 1571745 |
|  | STD |  |  | 8792.8 | 5433.0 |
|  | %RSD |  |  | 0.57 | 0.35 |

Precision of Sumatriptan and Naproxen

Table -8

Robustness:-

Robustness of Sumatriptan

| Proposed variations |  | USP Plate Count | USP Tailing |
| --- | --- | --- | --- |
| Variation in mobile phase composition | 10% less | 2396.9 | 1.8 |
|  | *Actual | 2183.4 | 1.6 |
|  | 10% more | 2218.9 | 1.5 |
| Variation in flow rate | 0.6ml/min | 2228.3 | 1.6 |
|  | 0.8ml/min | 2183.4 | 1.6 |
|  | 1.0ml/min | 2142.7 | 1.6 |

Table -9

Robustness of Naproxen

| Proposed variations |  | USP Plate Count | USP Tailing |
| --- | --- | --- | --- |
| Variation in mobile phase composition | 10% less | 2963.1 | 1.2 |
|  | *Actual | 2203.7 | 1.3 |
|  | 10% more | 2268.8 | 1.3 |
| Variation in flow rate | 0.6ml/min | 2519.4 | 1.3 |
|  | 0.8ml/min | 2203.7 | 1.3 |
|  | 1.0ml/min | 2474.4 | 1.3 |

Table -10

Ruggedness:

Ruggedness of Sumatriptan and Naproxen

|  | Retention Time of Naproxen | Retention Time of Sumatriptan | Area of Naproxen | Area of Sumatriptan |
| --- | --- | --- | --- | --- |
| Standard(80mcg) | 2.293 | 5.951 | 1976857 | 1990948 |
| Analyst(1)(80mcg) | 2.291 | 5.955 | 1971778 | 1993690 |
| Analyst(2)(80mcg) | 2.290 | 5.944 | 1970279 | 1995386 |
| Analyst(3)(80mcg) | 2.290 | 5.935 | 1979007 | 1992472 |
| Analyst(4)(80mcg) | 2.286 | 5.923 | 1970631 | 1993363 |
| AVG |  |  | 1973711 | 1993172 |
| STD |  |  | 3966.9 | 1631.1 |
| %RSD |  |  | 0.20 | 0.08 |

Figure 1:

chromatogram of Naproxen

Figure 2:

chromatogram of Sumatriptan

Figure 3:

Linearity:-

Linearity data of Naproxen

Figure 4:

Linearity data of  Sumatriptan:-

Figure 5:

Assay: - Sumatriptan and Naproxen

Standard

Sample

Figure 6:

Accuracy:-

Accuracy 50%

Accuracy 100%

Accuracy 150%

Figure 7:

PRECISSION:-

Injection 1

Injection 2

Injection 3

Injection 4

Injection 5

Figure 8:

System suitability:-

|  | Name | Retenction Time (min) | Area (µv*sec) | Height (µv) | USP Plate count | USP Tailing |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | Naproxen | 2.249 | 1531670 | 147652 | 2187.1 | 1.6 |
| 2 | Sumatriptan | 5.875 | 1573899 | 84946 | 2196.7 | 1.3 |

Figure 9:

Robustness

More organic solvent

Less organic solvent

More flow

Less flow

Figure 10:

Ruggedness:-

Injection 1

Injection 2

Injection 3

Injection 4

Injection 5

Figure 11:

Limit of detection:-

Sumatriptan

Naproxen

Figure 12:

Limit of quantification:-

Sumatriptan

Naproxen

Figure 13:

Structure of Sumatriptan

Figure 14:

Structure of Naproxen